## Summary
Optical Throughput of NFIRAOS. | [7 empty] | New LGS values JA 4-19-22 | Updated with Pyr WFS design, deleted MOR TWFS
TMT.AOS.TEC.10.059.REL07 | [8 empty] | Updated Silver coatings to FSS99-1500 data
Updated DM coatings with REQs from NWC DRD REL06
Typical Throughput
| LGS | VNW | I/Y-band | J-band | H-band | K-band |  | Comments | Technical Issues
| 0.589 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
Total throughput | 0.33775 | 0.35879 | 0.63098 | 0.74763 | 0.77219 | 0.76674
Telescope | 0.85737 | 0.91267 | 0.91267 | 0.91267 | 0.91267 | 0.91267 |  | REQ-OAD-1-1600 = 3-5% (Goal 2%) | 0.795/0.845/0.895 cf B.E. Mar. 14, 2008
Quantum Efficiency losses | 0.72 | 0.75 | [6 empty] | LGS:0.80/0.90/0.98 cf B.E. Mar. 14 & 17, 2008 NGS: Min QE >75%, Typ > 93% TMT.AOS.TEC.14.073
Total optical transmission | 0.54714 | 0.52416 | 0.69135 | 0.81917 | 0.84607 | 0.8401
NFIRAOS Contamination losses | 0.96315 | 0.95117 | 0.97528 | 0.97528 | 0.97528 | 0.97528 |  | loss of ~ 0.25% per optic | ~ 10 surfaces SCI, 15 LGS, 20 NGS
Science path optics | 0.80944 | 0.77674 | 0.70888 | 0.83993 | 0.86752 | 0.8614
Input Window (2) | 0.92237 | 0.86718 | 0.86718 | 0.9606 | 0.9606 | 0.98015 |  | Materion conceptual coating #2 see "coating curves" sheet. Concern for optic size, may not meet uniformity | Even with a double-sided coating, these windows have a 10:1 aspect ratio. Warping may occur. A 10+ layer design is expected.
OAP1 | 0.978 | 0.982 | 0.973 | 0.987 | 0.991 | 0.986 |  | protected silver FSS99-1600 | This technology is now owned by a different company. 98.5% may be optimistic for including contamination
DM11 | 0.971 | 0.975 | 0.981 | 0.98 | 0.989 | 0.986 |  | CILAS coating vendor proposed 2016
OAP2 | 0.978 | 0.982 | 0.973 | 0.987 | 0.991 | 0.986 |  | FSS99-1600
OAP3 | 0.978 | 0.982 | 0.973 | 0.987 | 0.991 | 0.986 |  | FSS99-1600
DM0 | 0.971 | 0.975 | 0.981 | 0.98 | 0.989 | 0.986 |  | CILAS coating vendor proposed 2016
Science B/S (reflect) | 0.995 | 0.995 |  |  |  |  | Reflect |  | front surface dichroic/ back surface AR could cause warping
Science B/S (transmit) |  |  | 0.974 | 0.972 | 0.966 | 0.97 | Transmit | Materion prototype coating, including AR
Science B/S (AR) | [6 empty] | Transmit | AR included above
OAP4-S |  |  | 0.973 | 0.987 | 0.991 | 0.986 |  | FSS99-1600
Selection fold |  |  | 0.973 | 0.987 | 0.991 | 0.986 |  | FSS99-1600
LGS optics | 0.70181 | 0.95
NGS/LGS B/S (transmit) |  | 0.95 |  |  |  |  | Transmit | GPI science B/S R>98% .7-.9 um; T>95% .94-2.4 um | Beamsplitter has a smaller wavelength range, but is still large and prone to warping. B/S will sag due to gravity
NGS/LGS B/S (reflect) | 0.98 | [5 empty] | Reflect
OAP4-L | 0.978 | [6 empty] | FSS99-1600 | Choose Science optics coating to allow oap to be used as a Science spare
Trombones | 0.98 | [6 empty] | 2x 1% dielectric coatings | dielectric V-coating should perform better than 99%, no issues
Periscope prism | 0.96 | [6 empty] | 1% dielectric x4, BK7 material
Collimating lens | 0.94 | [6 empty] | 1% per surface AR (3-lens coll)
LGS lenslets | 0.9 | [6 empty] | 0.88/0.92/0.96 cf B.E. Mar. 17, 2008
Relay lens | 0.92 | [6 empty] | 1% per surface AR (4-lens relay)
Visible Natural WFS optics |  | 0.7468
OAP4-V |  | 0.982 | [5 empty] | FSS99-1600
Large Fold mirror 1 |  | 0.99 | [5 empty] | protected silver optimized for 0.6-0.8 um
Large Fold mirror 2 |  | 0.99 | [5 empty] | protected silver optimized for 0.6-0.8 um | Could be changed to dielectric due to the narrower wavelength range. May not increase performance much. Silver can also be optimized for NIR
Fold mirror 3 (focus, centering) |  | 0.99 | [5 empty] | NIR dielectric | could also be silver
Lens 1 (relay doublet) |  | 0.97616 | [5 empty] | FPM2, NBH5 1% dielectric
T/T mirror (fast dithering) |  | 0.99 | [5 empty] | NIR dielectric | could also be silver
ADC 1 |  | 0.9748 | [5 empty] | TIM8, FPM2 1% dielectric
ADC 2 |  | 0.9748 | [5 empty] | TIM8, FPM2 1% dielectric
Packaging mirrors (1x) |  | 0.99 | [5 empty] | 1% dielectric per surface
Field stop |  | 1
Double Pyramid |  | 0.88769 | [5 empty] | NSL36, BAK4 1% dielectric
Lens 2 (relay triplet) |  | 0.97118 | [5 empty] | PBL35Y, NBH8, BSM2, 1% dielectric

## LGS only
Optical Throughput of NFIRAOS LGS
Revised September 9, 2009 JA
|  |  |  | BEST | WORST
DM concept
| LGS | Losses |  | LGS | LGS | Comments | LGS end photon requirement | [7 empty] | sodium return flux
Total throughput | 0.79206 | % |  | 0.31807 | 0.10073 |  | 900 ph per sub-ap per frame @ 800 Hz | [7 empty] | 500000 | ph/m^2/s/W
LGS optics
NGS/LGS B/S |  | 5 | Transmit |  |  | GPI science B/S R>98% .7-.9 um; T>95% .94-2.4 um | Energy of photon = hv |  |  |  | 900 | ph/.5mx.5m/800 Hz |  |  | uplink | 5 | W
NGS/LGS B/S | 0.98 | 2 | Reflect | 0.98 | 0.98 |  | h | 6.626e-34 | J s |  | 180000 | ph/m^2/s |  |  | flux | 2500000 | ph/m^2/s
OAP2b | 0.98 | 2 |  | 0.98 | 0.98 | protected silver R=98% 500-700 nm (denton) | v | 5.0934e+14 | 1/s |  | 18 | ph/cm^2/s |  |  |  | 250 | ph/cm^2/s
LGS Fold | 0.98 | 2 |  | 0.99 | 0.98 | 2% if metal coated, 1% if dielectric | Energy | 3.3749e-19 | J | per photon |  | ratio | 13.8889 | x flux available
Fixed Lens 1 | 0.96 | 4 |  | 0.96 | 0.96 | 2% per surface, negligable loss
Fixed Lens 2 | 0.96 | 4 |  | 0.96 | 0.96 | 2% per surface, negligable loss | 900 ph | 3.0374e-16 | J | per sub-ap per frame
trombone | 0.98 | 2 |  | 0.98 | 0.98 | 2x dielectric coatings | 2830 sub-ap | 8.5958e-13 | J | per full aperture per frame
pickoff mirror | 0.99 | 1 |  | 0.99 | 0.99 | 1x dielectric coatings
coll 1 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR | Frame rate | 800 | Hz
dm | 0.98 | 1 |  | 0.98 | 0.98 | 2% dielectric coating
coll 2 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR | Energy | 6.8766e-10 | J/s = W | full aperture per second
coll 3 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR
Fast Steer Mir | 0.99 | 1 |  | 0.99 | 0.99 | 1% dielectric | QE of det | 0.9 | %
LGS lenslets | 0.9 | 10 |  | 0.96 | 0.88 | 0.88/0.92/0.96 cf B.E. Mar. 17, 2008 | Energy | 7.6407e-10 | W
WFS relay | 0.96 | 4 |  | 0.96 | 0.96 | 2 lenses, 1% per surface | Opt TP | 0.6 | %
LGS CCD | 0.9 | 10 |  | 0.98 | 0.8 | 0.80/0.90/0.98 cf B.E. Mar. 14 & 17, 2008 | Energy | 1.2735e-09 | W
[7 empty] | Other loss | 0.9 | %
[7 empty] | Energy | 1.4149e-09 | W
[8 empty] | 1.4149 | uW
Revised February 11, 2010
DM concept with 4 OAP design
| LGS | Losses |  | LGS | LGS | Comments
Total throughput | 0.84226 | % |  | 0.47236 | 0.20194
LGS optics
NGS/LGS B/S |  | 5 | Transmit |  |  | GPI science B/S R>98% .7-.9 um; T>95% .94-2.4 um
NGS/LGS B/S | 0.98 | 2 | Reflect | 0.98 | 0.98
OAP2b | 0.98 | 2 |  | 0.98 | 0.98 | protected silver R=98% 500-700 nm (denton)
LGS Fold |  |  |  | 0.99 |  | removed
Fixed Lens 1 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface, negligable loss
Fixed Lens 2 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface, negligable loss
trombone | 0.98 | 2 |  | 0.98 | 0.98 | 2x dielectric coatings
pickoff mirror | 0.99 | 1 |  | 0.99 | 0.99 | 1x dielectric coatings
coll 1 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR
dm | 0.98 | 1 |  | 0.98 | 0.98 | 2% dielectric coating
coll 2 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR
coll 3 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR
Fast Steer Mir | 0.99 | 1 |  | 0.99 | 0.99 | 1% dielectric
LGS lenslets | 0.9 | 10 |  | 0.96 | 0.88 | 0.88/0.92/0.96 cf B.E. Mar. 17, 2008
WFS relay | 0.96 | 4 |  | 0.96 | 0.96 | 2 lenses, 1% per surface
LGS CCD | 0.9 | 10 |  | 0.98 | 0.8 | 0.80/0.90/0.98 cf B.E. Mar. 14 & 17, 2008
Revised September 21, 2010
cold concept with 4 OAP design
| LGS | Losses |  | LGS | LGS | Comments
Total throughput | 0.93178 | % |  | 0.6465 | 0.36566
LGS optics
NGS/LGS B/S |  | 5 | Transmit |  |  | GPI science B/S R>98% .7-.9 um; T>95% .94-2.4 um
NGS/LGS B/S | 0.98 | 2 | Reflect | 0.98 | 0.98
OAP2b | 0.98 | 2 |  | 0.98 | 0.98 | protected silver R=98% 500-700 nm (denton)
LGS Fold |  |  |  | 0.99 |  | removed
Fixed Lens 1 |  | 2 |  |  |  | removed
Fixed Lens 2 |  | 2 |  |  |  | removed
trombone | 0.98 | 2 |  | 0.98 | 0.98 | 2x dielectric coatings
pickoff mirror | 0.99 | 1 |  | 0.99 | 0.99 | 1x dielectric coatings
coll 1 |  | 2 |  |  |  | removed
dm |  | 1 |  |  |  | removed
coll 2 |  | 2 |  |  |  | removed
coll 3 | 0.98 | 2 |  | 0.98 | 0.98 | 1% per surface AR
Fast Steer Mir | 0.99 | 1 |  | 0.99 | 0.99 | 1% dielectric
LGS lenslets | 0.9 | 10 |  | 0.96 | 0.88 | 0.88/0.92/0.96 cf B.E. Mar. 17, 2008
WFS relay | 0.96 | 4 |  | 0.96 | 0.96 | 2 lenses, 1% per surface
LGS CCD | 0.9 | 10 |  | 0.98 | 0.8 | 0.80/0.90/0.98 cf B.E. Mar. 14 & 17, 2008
Revised April 26, 2011
cold concept with 4 OAP design
| LGS | Losses |  | LGS | LGS | Comments
Total throughput | 0.91276 | % |  | 0.79981 | 0.5985
LGS optics
NGS/LGS B/S |  | 5 | Transmit |  |  | GPI science B/S R>98% .7-.9 um; T>95% .94-2.4 um
NGS/LGS B/S | 0.98 | 2 | Reflect | 0.98 | 0.98
OAP2b | 0.98 | 2 |  | 0.98 | 0.98 | protected silver R=98% 500-700 nm (denton)
trombone | 0.98 | 2 |  | 0.98 | 0.98 | 2x dielectric coatings
pickoff mirror | 0.99 | 1 |  | 0.99 | 0.99 | 1x dielectric coatings
coll 3 | 0.96 | 4 |  | 0.96 | 0.96 | 1% per surface AR (2-lens design)
Fast Steer Mir | 0.99 | 1 |  | 0.99 | 0.99 | removed
LGS lenslets | 0.9 | 10 |  | 0.96 | 0.88 | 0.88/0.92/0.96 cf B.E. Mar. 17, 2008
WFS relay | 0.96 | 4 |  | 0.96 | 0.96 | removed
LGS CCD | 0.9 | 10 |  | 0.98 | 0.8 | 0.80/0.90/0.98 cf B.E. Mar. 14 & 17, 2008

## Coating curves
AR for Entrance Windows
Materion
Prototype AR used in conjunction
with dichroic (700-2400 nm).
They expect very minor
improvements can be
made in the AR coating,
especially when extending to
500 nm on the blue side.
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band | [5 empty] | LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45 | [5 empty] | 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| n/a | n/a | n/a | 98 | 98 | 98 | 99 | [5 empty] | 98 | 96.5 | 96.5 | 96.5 | 99 | 99 | 99.5
Evaporated Coatings
They have coated small parts
for 600-2400 nm, but need
development time for a large
optic.
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 98.5 | 98.5 | 98.5 | 98.3 | 98.6 | 98.6 | 98.6
Paolo's model
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 99.5 | 96.5 | 96.5 | 98 | 99 | 99 | 99
Protected silver for mirrors
Quantum makes the old Denton.
The overcoat can be adjusted to
peak performance at a given
wavelength
| red curve: | [6 empty] | green curve:
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band | LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45 | 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 97.3 | 97.5 | 97.5 | 96.7 | 98.7 | 99 | 98 | 98.5 | 97.5 | 97.5 | 97.2 | 97.2 | 98.1 | 98.7
Cilas uses an unknown vendor
They have a proposed coating
used for GTC and space qual.
This was part of the 2013
roadmap presentation
In 2011 they sent a typical
coating curve that performs
better than the GTC
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band | [5 empty] | LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45 | [5 empty] | 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 94.2 | 94.1 | 94.1 | 95 | 97 | 98 | 98.3 |  |  | REQ-NWC-1270, NWC-1470 |  |  | 96 | n/a | 96 | 97.5 | 98.5 | 99 | 99
Dichroic Beamsplitter
Materion built a prototype b/s
with a dichroic and AR coating
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 99.5 | 99.5 | 99.5 | 97.4 | 97.2 | 96.6 | 97
GPI beamsplitter was coated
at Coastline Optics. They don't
have capabilities for our size.
The design is very similar to
what Materion presented.
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| 99.5 | 99 | 99 | n/a | 98.2 | 98.2 | 90
Zygo also made a conceptual
coating design
We don't have the reflected data
| LGS | MOR | NGS | I/Y-band | J-band | H-band | K-band
| 0.589 | 0.6-0.75 | 0.6-0.75 | 0.8-1.0 | 1.1-1.4 | 1.45-1.85 | 1.9-2.45
| n/a | n/a | n/a | 95.6 | 96 | 96.2 | 97
Detectors
E2V CCD 201
baseline for MOR TWFS